## Slide 1: Magnetically Enhanced Microflow Cytometer for Bead-based Immunoaffinity Measurements in Blood Samples
Master Thesis Presentation
Johann A. Brenner
Munich, 31.03.2021
Notes: Quick Note:
I don’t see your video very well during the presentation  Please unmute and interrupt me, especially if I lose you at some point which is hopefully not going to be the case

## Slide 2: Outline
Theory
  Immuno-Affinity Measurements in the Magnetic Flow Cytometer
  Surface Chemistry Methods
Results
    Simulation
      Virtual Prototyping of Cell Signals
      Calculus Estimation of the Force-Equilibrium in a Microchannel
    Experimental
      Magnetic Concentration Measurements
      Development of a Differential Counting Setup
      Established Reference Model for Surface Marker Expression Densities
      Affinity-based Concentration Measurements in the Magnetic Flow Cytometer
Conclusion

## Slide 3: Theory
Magnetic bead rolling on the floor of a microchannel
Biofunctionalized Channel
Force-Equilibrium:
  Hydrodynamics
  Volume Forces
  Electro-Magnetism
   Change in Velocity or Adhesion
Immuno-Affinity Measurements in the Magnetic Flow Cytometer
Notes: Quick Note: Beads = Reference Model Cells during this thesis but are used equivalently

## Slide 4: Theory
Immuno-Affinity Measurements in the Magnetic Flow Cytometer
Notes: Please excuse the unscientificy animation, but it illustrates the principle.

## Slide 5: Theory
Immuno-Affinity Measurements in the Magnetic Flow Cytometer

## Slide 6: Theory
Immuno-Affinity Measurements in the Magnetic Flow Cytometer
2.5 mm
Helou et al. (2013) Lab on a Chip, 13(6)

## Slide 7: Theory
Immuno-Affinity Measurements in the Magnetic Flow Cytometer

## Slide 8: Theory
  Surface Chemistry Methods – Covalent Modification
Oxidation
Silanization
Polymer Coating
Protein
Attachment

## Slide 9: Theory
  Surface Chemistry Methods – Biotin-Streptavidin
Biotin-Streptavidin Interaction
One of the strongest known non-covalent bonds
Affinity: Kd = 10-13 - 10-15
Release force: 100 – 400 pN
  Adhesion
Sedlak et al. (2020) Sci Adv, 6(13)

## Slide 10: Theory
  Surface Chemistry Methods – Biotin-Streptavidin

## Slide 11: Results
Simulation – Virtual Prototyping of Cell Signals
MNP-Coverage Evaluation
Aggregate Angle Evaluation
Notes: Simulation of cells which are labelled magnetically and flowing over the sensor bridge
Epitope density should be modelled
Left: Signal dependency of coverage
  Compared to a reference dipole indicated by the blue circles  if all momentum is located in the center
Right: Influence of cell aggregates on the signal shape and their deviation from the reference dipole

## Slide 12: Results
Low Coverage:
  Single Magnetic Momenta cause field inhomogenity
Signal deterioration for larger MNP diameters
Simulation – Virtual Prototyping of Cell Signals
Notes: 641 600 MNP with 20 nm
81 MNP with 2 μm
Left:
  Low Coverage:
    Single Magnetic Momenta cause field inhomogenity
  magnetic momentum per MNP as it is dependent on the volume - thus r3.
    Larger Particles cause more likely inhomogeneities
Right:
  Maximum correlation shows also a significant deterioration for larger cell diameter
  It is assumed, that the correlation error becomes constant at great Particles
  Reason
  Particles have lower surface area  Less beads on surface  Inhomogenities --> microparticles have a higher influence,

## Slide 13: Results
Simulation – Virtual Prototyping of Cell Signals

## Slide 14: Results
High correlation in the medium coverage regime due to inhomgeneous field
Low yaw and high coverage  Two dipoles in superposition
High yaw and coverage  Aggregate cannot be discriminated by the sensor
Simulation – Virtual Prototyping of Cell Signals
Notes: two 200 nm MNP-covered cells
Interestingly the error for 15° and 30° is largest
Low coverage: Field inhomogeneous, but not influenced by two spheres
For low yaw angle and high coverages
  the aggregate’s signal reflects rather two single dipoles in superposition than one quite homogeneous dipole.
High yaw angle, and high coverages:
  Two beads parallel to the large sensor side cannot be discriminated  similar correlation

## Slide 15: Results
Simulation – Calculus Estimation of the Force-Equilibrium in a Microchannel
Flow Field in a Particle
Shear Field in a Particle
Notes: Framework has been developed to analyze the force equilibrium inside the channel and on the rolling bead
Flow induces rolling motion  Top is actuated more
However, high shear acts at the wall and thus the potential biological bonds.

## Slide 16: Results
Simulation – Calculus Estimation of the Force-Equilibrium in a Microchannel
biotin
neutravidin
bead
z
Adopted from Wu and Voldman (2020) Biosensors and Bioelectronics, 154, 112070.
Notes: By the analytical modelling of many forces, the dominant actors were identified
Stoke’s Drag Force, Magnetophoretic Force and Protein Interaction
fluid forces 500 pN in the flow direction (463.5 for 80µlmin)
Perpendicular
Magnetophoretic Structures:
  for a 8 micron bead: 5 nN (5 kT/m) in the nickel-iron structures
Protein interaction:
  8 micrometer bead with 1 µmol/g surface charges
  177500 per bead
  Bead approach: 30 nm

## Slide 17: Results
Simulation – Calculus Estimation of the Force-Equilibrium in a Microchannel
biotin
neutravidin
bead
z
Adopted from Wu and Voldman (2020) Biosensors and Bioelectronics, 154, 112070.
1.5 µm2
~ 1300 Proteins
Notes: By the analytical modelling of many forces, the dominant actors were identified
Stoke’s Drag Force, Magnetophoretic Force and Protein Interaction
fluid forces 500 pN in the flow direction (463.5 for 80µlmin)
Perpendicular
Magnetophoretic Structures:
  for a 8 micron bead: 5 nN (5 kT/m) in the nickel-iron structures
Protein interaction:
  8 micrometer bead with 1 µmol/g surface charges
  177500 per bead
  Bead approach: 30 nm  Interaction area of 1.5 µm^2 or 0.75% of the particle
  1300 proteins bond potentially
  Different Protein relase forces lead to 5-200 nN adhesion
rolling resistance factor for a perfectly elastic surface from the car and bike industry
  Project perpendicular forces in flow direction

## Slide 18: Theory
  Surface Chemistry Methods – Biotin-Streptavidin
Predominant Forces
Stoke‘s Drag: 460 pN
Dipolar Force: 5 nN
Protein Force: 5-200 nN
Unfunctionalized Bead
150 pN Drag
Bio-Functionalized Bead
Adhesion caused by 16 - 500 molecules of neutravidin
Notes: Assuming an unfunctionalized bead this lead to 150 pN in flow
Interaction of 16 to 500 neutravidin molecules (respective release force)

## Slide 19: Results
Bead stocks characterized in the flow cytometer precisely
Empirical correction factors to account for bead losses in interconnections and flow profile:
    Experimental – Magnetic Concentration Measurements
Notes: Measure adhesion  Concentration Measurements neccessary
Measured counts lower than in sample  losses in syringe, tubing, flow profile
Correction factor:
  With previously known concentration  Cconst  Normalized to a standard procedure
  If velocity is determined robustly  Cvelocity  Fraction of the Particle velocity from the flow profile.

## Slide 21: Results
    Experimental – Development of a Differential Counting Setup
Notes: To get rid of the absolute counting and the related errors,  relative/differential counting setup was built from existing components
In plane integration not possible  window of perpendicular field lines is too narrow
Non perpendicular components  Exactly measurement direction
Stacking of the two PCB-carriers with spacers in between
drilled hole in the top to interconnect

## Slide 22: Results
    Experimental – Development of a Differential Counting Setup
Notes: To get rid of the absolute counting and the related errors,  relative/differential counting setup was built from existing components
In plane integration not possible  window of perpendicular field lines for the magnetophoresis is too narrow
Non perpendicular components  Exactly measurement direction
Stacking of the two PCB-carriers with spacers in between
drilled hole in the top to interconnect
However, sensitivities are not really promising, but sufficient for beads
Due to the height difference, the gradients are also different in magnitude  Magnetic field  r^3
What does that mean?
  Low flow rate: Slow beads, perfectly pulled down on top and measured, Adhesion at the bottom  Force equilibrium dominated by magnetic force
  High Flow rate: Beads to fast to get pulled down in the low gradient on top
Optimal Flow rate determined by the quotient  Optimum at 160 µl
Notwithstanding, magnitude of measured beads is too low  Omitting this idea

## Slide 23: Results
    Experimental – Development of a Differential Counting Setup

## Slide 24: Results
Proof of concept for the simulations in the magnetic flow cytometer
Epitope Density on a cell has to be modulated
Development of a methodology to coat beads with highly defined surface markers
Universally applicable to proteins
 Carbodiimide Chemistry
    Experimental – Established Reference Model for Surface Marker Expression Densities

## Slide 25
Notes: Carboxy left, Amine in center, Antibody on carboxylate right

## Slide 26
Notes: Half life of carboxylate beads  1.4 days  Can probably be attributed to improper washing
 Saw bead disintegration  Something happens

## Slide 27
Concenctration measurement of magnetized polystyrene beads
    Experimental – Established Reference Model for Surface Marker Expression Densities
Results
Notes: Non-magnetic polystyrene beads functionalized by previous protocol
Magnetized with 50nm MNPs
Weakly, but significant difference between functionalization densities visible
Decrease in relative counts with less mnp coverage
  Get not pulled to bottom
Velocity Distribution of Fully and partly coated beads shows equal mean
 no difference in rolling attributed to more magnetic force measurable

## Slide 28: Results
    Experimental – Affinity-based Concentration Measurements in the Magnetic Flow Cytometer
Biotinylated Beads
Plain Beads
Notes: These biotinylated beads were pumped over a covalently functionalized surface
Microscope images showed promising results
PDMS works best
SiN works worst
Worse Glass attributed to all measurements in the beginning of the thesis
Not sure if covalent or not  No adequate method in reach

## Slide 29: Results
Experimental – Affinity-based Concentration Measurements in the
Magnetic Flow Cytometer
Notes: Titration of surface functionalization:
Neutravidin Titrated :
  Again Bead Density evaluation
  Fluorescence Microscopy Evaluation  Channel imaged and referenced to a blank
However:
  covalent modification is time and ressource consuming
  Unspecific Physisorption is not much worse  Will be used in further work

## Slide 30
Plain Particles
Biotinylated Particles
Notes: Sensor chip functionalized with neutravidin after the first sensor

## Slide 31
    Experimental – Affinity-based Concentration Measurements in the Magnetic Flow Cytometer
Results
Notes: Affinity-base Concenctrations measured magnetically:
- Previously titrated surface coverage of beads now used in the affinity measurements
- Not only very large dynamic range, but also a coverage dependent behavior visible
- Properly functionalized sensor
 Very robust and constant counting
 adhesion of beads
 Initial rise attributed to filling of magnetophoretic structures

## Slide 32: Conclusion
Established hydrodynamic and magnetic simulations for physical phenomena of rolling cells
Benchmarked absolute and differential concentration measurements for the magnetic flow cytometer
Developed a reference system for the variation of the surface receptor density
Demonstrated an affinity-dependent assay in the magnetic flow cytometer
Outcomes of this thesis
Notes: Physical phenomena of rolling cells with a bio-functionalized surface under laminar flow conditions were be simulated from a hydrodynamic and magnetic point of view
  synergy of fluid dynamics, inertia, viscous, and body forces was illuminated in an analytical force equilibrium
  numeric magnetic field simulation of MNP-laden spheres was correlated to dipole signals in the magnetic flow cytometer get further insights into signal patterns and possible parameter estimations
Absolute and differential concentration measurements were benchmarked for the magnetic flow cytometer
  As a result, empirical correction factors, to account for losses in syringe, connectors, and the parabolic flow profile, were determined
  To overcome these limitations  a relative counting apparatus has been engineered from two serially connected magnetic sensors.
  However, the complexity and tentativeness of the differential system requires further developments beyond this thesis
A reference system for the variation of surface receptor density was established and subsequently evaluated
  With the developed methods, several proteins were attached to the surface of microbeads covalently
  Protocols optimized to preserve a reliable dose titration of biotin and antibodies for the coating density
  Furthermore, the a magnetic measurement of weakly magnetized cells was evaluated in a pretest by MNP-coating the functionalized microbeads.
an affinity-based concentration assay was presented which reveals promising results for the magnetic measurement of biomarker concentrations or single cell surface proteins in the future
  beads were immobilized in a laminar flow according to their respective biomolecular bond strength
  Modified channels via physisorption and covalent attachment on various surfaces such as PDMS, Si3N4, and glass
  right now physisorption is the favorized technique although it dismantles swiftly.
  With these biologically augmented microfluidics, magnetic measurements showed a clear separation from fully- to sparsely-coated beads
  Even, a dependency on the biotinylation-degree of microbeads has been measured with strong statistical support.

## Slide 33: Thank you, LBE!

## Slide 34: Magnetic Field Computations

## Slide 35: Error Sources in the Magnetic Flow Cytometer
Losses in Interconnections
Systematic Error in Velocity Measurements at High Concentrations

## Slide 36: Affinity-dependent Bead Velocity
Magnetically measured bead velocity of biotinlyated bead is statistically higher
Probably attributed to a saltatory movement after adhesion to the channel (B,C)
After adhesion, a wake forms behind the bead thus causes increasing pressure on the back

## Slide 37: Comparison of Empirically Determined Correction Factors
Errors in Affinity-Based Measurements

## Slide 38: Error Sources Affinity-Dependent Concentration Measurement
Inital 100 µl:
  magnetophoretic structures are filled
  Steep rise
After 700µl:
  Surface is saturated with biotinylated beads
  Burst-Like Measurement of beads
  After 800µl:
    Empty Syringe and Sedimentation Effects play a role
     Less beads than expected are measured

## Slide 39: Hysteresis Calibration of the Sensor

## Slide 40: Micromod BNF-Dextran

## Slide 41: Assessment of Physisorbed Biofunctionalization
Platereader Experiment with AlOx and SiN surfaces
Coverage with fluorescent streptavidin
Washing by 200 µl of buffer with a micropipette
No stability over multiple washing steps
SiN superior to AlOx

## Slide 42: Fluorescent Evaluation of the Biofunctionalization
Image Analysis: Integration over vertical Axis
Tool: ImageJ

## Slide 43: Surface Activation with Plasma
Difference to Piranha not measrurable in contact angle, but in functionalization operation

## Slide 44: 8µm Beads coated with streptavidin-redF-Dextran
Maximum Biotin Coverage
Blank
Approx 16% Coverage
Unspecific Binding

## Slide 45: Anti-Biotin-PE Assessment Not Working

## Slide 46: Wafer - Spincoating
17.07 3“ SU8-3050 @ 3000 rpm written same day
23.07 3“ SU8-3050 @ 3000 rpm written after one week
27.07 3“ SU8-3050 @ 3000 rpm written same day
29.07 3“ SU8-3025 @ 3000 rpm written two days after
4“ desired 150µm Layer
~ 98 µm
~ 94 µm
~ 53 µm
~ 44 µm
~ 115µm
Water Desorption 120min@160°C
Cool to RT
3ml SU8 (1ml/inch Wafer diameter)
Softbake 1min@65°C 15min@95°C(50µm) 10min@95°C(25µm)
Laser writing / Flood exposure
Spincoat 5s@500rpm(100rpm/s) 30s@3000rpm(300rpm/s)
Clean with Aceton 30s@7000rpm
Post-Exposure-Bake 1min@65°C 5°C@95°C
Develop mrDev 600 6-10min with shaking Optionally: in sonication Bath
Cool to RT

## Slide 47: Wafer – Laserwriting
Parameter | Value
Spot size | 0.01 mm
Overlap | 80%
Modulation | 100%
Z-Position | -6.782
velocity | 40 mm/s (omitted lines in dialog)