## Slide 1: Magnetically Enhanced Microflow Cytometer for Bead-based Immunoaffinity Measurements in Blood Samples
Master Thesis Presentation
Johann A. Brenner
Munich, 29.03.2021

## Slide 2: Outline
Theory
  Immuno-Affinity Measurements in the Magnetic Flow Cytometer
  Surface Chemistry Methods
Results
    Simulation
      Virtual Prototyping of Cell Signals
      Calculus Estimation of the Force-Equilibrium in a Microchannel
    Experimental
      Magnetic Concentration Measurements
      Development of a Differential Counting Setup
      Established Reference Model for Surface Marker Expression Densities
      Affinity-based Concentration Measurements in the Magnetic Flow Cytometer
Conclusion

## Slide 3: Theory
Magnetic bead rolling on the floor of a microchannel
Biofunctionalized Channel
Force-Equilibrium:
  Hydrodynamics
  Volume Forces
  Electro-Magnetism
   Change in Velocity or Adhesion
Immuno-Affinity Measurements in the Magnetic Flow Cytometer

## Slide 4: Theory
Immuno-Affinity Measurements in the Magnetic Flow Cytometer

## Slide 5: Theory
Immuno-Affinity Measurements in the Magnetic Flow Cytometer

## Slide 6: Theory
Immuno-Affinity Measurements in the Magnetic Flow Cytometer

## Slide 7: Theory
Immuno-Affinity Measurements in the Magnetic Flow Cytometer
2.5 mm
Helou et al. (2013) Lab on a Chip, 13(6)

## Slide 8: Theory
Immuno-Affinity Measurements in the Magnetic Flow Cytometer

## Slide 9: Theory
  Surface Chemistry Methods – Covalent Modification
Oxidation
Silanization
Polymer Coating
Protein
Attachment

## Slide 10: Theory
  Surface Chemistry Methods – Particles

## Slide 11: Theory
  Surface Chemistry Methods – Biotin-Streptavidin
Biotin-Streptavidin Interaction
One of the strongest known non-covalent bonds
Affinity: 10-13 - 10-15
Release force: 100 – 400 pN
  Adhesion
Sedlak et al. (2020) Sci Adv, 6(13)

## Slide 12: Theory
  Surface Chemistry Methods – Biotin-Streptavidin

## Slide 13: Results
Simulation – Virtual Prototyping of Cell Signals
MNP-Coverage Evaluation
Aggregate Angle Evaluation
Notes: Simulation of cells which are labelled magnetically and flowing over the sensor bridge
Epitope density should be modelled
Left: Signal dependency of coverage
  Compared to a reference dipole indicated by the blue circles  if all momentum is located in the center
Right: Influence of cell aggregates on the signal shape and their deviation from the reference dipole

## Slide 14: Results
Low Coverage:
  Single Magnetic Momenta cause field inhomogenity
Signal deterioration for larger cell diameters
Simulation – Virtual Prototyping of Cell Signals
Notes: 641 600 MNP with 20 nm
81 MNP with 2 μm
Left:
  Low Coverage:
    Single Magnetic Momenta cause field inhomogenity
  magnetic momentum per MNP as it is dependent on the volume - thus r3.
    Larger Particles cause more likely inhomogeneities
Right:
  Maximum correlation shows also a significant deterioration for larger cell diameter
  It is assumed, that the correlation error becomes constant at great Particles

## Slide 15: Results
Simulation – Virtual Prototyping of Cell Signals

## Slide 16: Results
High correlation in the medium coverage regime due to inhomgeneous field
Low yaw and high coverage  Two dipoles in superposition
High yaw and coverage  Aggregate cannot be discriminated by the sensor
Simulation – Virtual Prototyping of Cell Signals
Notes: two 200 nm MNP-covered cells
Interestingly the error for 15° and 30° is largest
Low coverage: Field inhomogeneous, but not influenced by two spheres
For low yaw angle and high coverages
  the aggregate’s signal reflects rather two single dipoles in superposition than one quite homogeneous dipole.
High yaw angle, and high coverages:
  Two beads parallel to the large sensor side cannot be discriminated  similar correlation

## Slide 17: Results
Simulation – Calculus Estimation of the Force-Equilibrium in a Microchannel
Flow Field in a Particle
Shear Field in a Particle
Notes: Framework has been developed to analyze the force equilibrium inside the channel and on the rolling bead
Flow induces rolling motion  Top is actuated more
However, high shear acts at the wall and thus the potential biological bonds.

## Slide 18: Results
Simulation – Calculus Estimation of the Force-Equilibrium in a Microchannel

## Slide 19: Results
    Experimental – Magnetic Concentration Measurements

## Slide 21: Results
    Experimental – Development of a Differential Counting Setup

## Slide 22: Results
    Experimental – Development of a Differential Counting Setup

## Slide 23: Results
    Experimental – Established Reference Model for Surface Marker Expression Densities

## Slide 25
Ocean Nanotech MNPs
    Experimental – Established Reference Model for Surface Marker Expression Densities
Results

## Slide 26: Results
    Experimental – Affinity-based Concentration Measurements in the Magnetic Flow Cytometer

## Slide 27: Results
Experimental – Affinity-based Concentration Measurements in the
Magnetic Flow Cytometer

## Slide 28
Plain Particles
Biotinylated Particles

## Slide 29
    Experimental – Affinity-based Concentration Measurements in the Magnetic Flow Cytometer
Results

## Slide 30: Conclusion

## Slide 31: Thank you!

## Slide 32: Magnetic Field Computations

## Slide 33: Error Sources in the Magnetic Flow Cytometer

## Slide 34: Affinity-dependent Bead Velocity
Magnetically measured bead velocity of biotinlyated bead is statistically higher
Probably attributed to a saltatory movement after adhesion to the channel (B,C)
After adhesion, a wake forms behind the bead thus causes increasing pressure on the back

## Slide 35: Comparison of Empirically Determined Correction Factors

## Slide 36: Error Sources Affinity-Dependent Concentration Measurement
Inital 100 µl:
  magnetophoretic structures are filled
  Steep rise
After 700µl:
  Surface is saturated with biotinylated beads
  Burst-Like Measurement of beads
  After 800µl:
    Empty Syringe and Sedimentation Effects play a role
     Less beads than expected are measured

## Slide 37: Hysteresis Calibration of the Sensor

## Slide 38: Ocean Nanotech MNPs

## Slide 39: Micromod BNF-Dextran

## Slide 40: Assessment of Physisorbed Biofunctionalization
Platereader Experiment with AlOx and SiN surfaces
Coverage with fluorescent streptavidin
Washing by 200 µl of buffer with a micropipette
No stability over multiple washing steps
SiN superior to AlOx

## Slide 41: Anti-Biotin-PE Assessment

## Slide 42: Fluorescent Evaluation of the Biofunctionalization

## Slide 43: Surface Activation with Plasma
Difference to Piranha not measrurable in contact angle, but in functionalization operation

## Slide 44: 8µm Beads coated with streptavidin-redF-Dextran
Maximum Biotin Coverage
Blank
Approx 16% Coverage

## Slide 45: 02.06. Biotinylation non-magnetic 8µm
Biotinylated, non-magnetic 8µm Polystyrol
Label: Streptavidin Atto-488
MES
TRIS
TAE
High
(1,dark green in FACS)
Blank
(5,light green in FACS)
µ=52, =11
µ=42, =19
µ=70, =14
µ=40, =9
µ=24, =0
µ=54, =16

## Slide 46: Wafer - Spincoating
17.07 3“ SU8-3050 @ 3000 rpm written same day
23.07 3“ SU8-3050 @ 3000 rpm written after one week
27.07 3“ SU8-3050 @ 3000 rpm written same day
29.07 3“ SU8-3025 @ 3000 rpm written two days after
4“ desired 150µm Layer
~ 98 µm
~ 94 µm
~ 53 µm
~ 44 µm
~ 115µm
Water Desorption 120min@160°C
Cool to RT
3ml SU8 (1ml/inch Wafer diameter)
Softbake 1min@65°C 15min@95°C(50µm) 10min@95°C(25µm)
Laser writing / Flood exposure
Spincoat 5s@500rpm(100rpm/s) 30s@3000rpm(300rpm/s)
Clean with Aceton 30s@7000rpm
Post-Exposure-Bake 1min@65°C 5°C@95°C
Develop mrDev 600 6-10min with shaking Optionally: in sonication Bath
Cool to RT

## Slide 47: Wafer – Laserwriting
Parameter | Value
Spot size | 0.01 mm
Overlap | 80%
Modulation | 100%
Z-Position | -6.782
velocity | 40 mm/s (omitted lines in dialog)